Title: Tensile tests of interface between bone and plasma-sprayed HA coating-titanium implant

Author(s): Lin H, Xu HC, Zhang XD, de Groot K

Source: JOURNAL OF BIOMEDICAL MATERIALS RESEARCH 43 (2): 113-122 SUM 1998

Document Type: Article

Language: English

Cited References: 18      Times Cited: 9

Abstract: Tensile strength testing of the interface between bone and a plasma-sprayed hydroxyapatite (HA) coating-Ti-6Al-4V (implant I) and an HA coating-Ti (implant II) was performed. The bone-implant interface and tensile failure mode were evaluated by light microscopy (LM), SEM-energy dispersive X-ray analysis, and backscattered electron imaging. The crystallinity of the HA coating of implant I and implant II was 62 and 40%, respectively. Eleven semidisk implants of each type were inserted in the left tibia of 22 rabbits in each period. After 2, 4, 8, and 16 weeks, the tensile strength of the bone-implant I was 0.86, 1.32, 1.10, and 0.92 MPa, respectively; the value of the bone-implant II was 0.66, 0.92, 0.84, and 1.12 MPa, respectively. No significant difference was found in tensile strength between implant I and implant II, and between the data of different periods. LM and SEM revealed that the two types of coatings behave the same in bone. Failure after the tensile test mainly occurred at the bony tissue at 2 and 4 weeks, inside the HA coating layer at 8 weeks, and at the HA coating-metal substrate interface at 16 weeks. Degradation and delamination of the HA coating continued with time, resulting in the interfacial tensile strength not increasing with time. Thus, we concluded that differences between the two implants, mainly the difference in crystallinity of the HA coating (62 vs. 40%), have no significant influence on bone bonding ability, bone bonding strength, tensile failure mode, and degradation of the coating. (C) 1998 John Wiley & Sons, Inc.

Author Keywords: plasma-sprayed HA coating; HA coating metal substrate; tensile strength; degradation; bone bonding

KeyWords Plus: COATED IMPLANTS; HYDROXYAPATITE; STRENGTH

Addresses: de Groot K (reprint author), Leiden Univ, Biomat Res Grp, Prof Bronkhorstlaan 10, Bilthoven, NL-3723 MB Netherlands
Leiden Univ, Biomat Res Grp, Bilthoven, NL-3723 MB Netherlands
Beijing Med Univ, Stomatol Coll, Dent Mat Lab, Beijing, 100081 Peoples R China
Sichuan Univ, Anal Testing Ctr, Chengdu, 610064 Peoples R China

Publisher: JOHN WILEY & SONS INC, 605 THIRD AVE, NEW YORK, NY 10158-0012 USA

Subject Category: ENGINEERING, BIOMEDICAL; MATERIALS SCIENCE, BIOMATERIALS; MATERIALS SCIENCE, MULTIDISCIPLINARY

IDS Number: ZR307

ISSN: 0021-9304